REAGEN™

Salbutamol ELISA Test Kit

RND99038

### Product Description

REAGEN™ Salbutamol ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Salbutamol in feed, fish, honey, meat (beef, chicken and pork), milk, milk powder, serum, and urine. The unique features of the kit are:

- Rapid (10 - 30 minutes) extraction method for various samples with high recovery (75 - 115%).

- High sensitivity (0.05 ng/g or ppb) and low detection limit (0.025 ng/g or ppb for meat/tissue, 0.25 ng/g or ppb for urine and 1 ng/g or ppb for animal feed).

- High reproducibility.

- A quick ELISA assay.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Salbutamol ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Salbutamol -coated Plate | 1 x 96-well (8 wells x 12 strips) | 2-8C |
| Salbutamol Standards: Negative control (white cap tube) 0.05 ng/mL (yellow cap tube) 0.15 ng/mL (orange cap tube) 0.5 ng/mL (pink cap tube) 1.5 ng/mL (purple cap tube) 4.5 ng/mL (blue cap tube) 10 ng/mL (spiking, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C |
| Salbutamol Antibody #1 | 12 mL | 2-8C |
| 100X HRP-Conjugated Antibody #2 | 250 L |  |
| Antibody #2 Diluent ** | 20 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| Stop Buffer ** | 14 mL |  |
| TMB Substrate ** | 12 mL |  |
| Sample Balance Buffer I ** Sample Balance Buffer II(powder) 10X PBS ** 20X Sample Extraction Buffer ** Liver Extraction Buffer (Optional) Beta-Agonist Clean Up Buffer I (Optional) Beta-Agonist Clean Up Buffer II (Optional) Beta-Agonist Clean Up Buffer III (Optional) | 10 mL 62g 25 mL 10 mL 10ml 5ml 5ml 28ml |  |

- If you are not planning to use the kit for over 1 months, store Salbutamol Antibody #1 and 100X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Feed | 1 |
| Honey | 1 |
| Meat/Liver | 0.1 |
| Swine/Bovine Liver and Kidney | 0.1 |
| Milk/serum | 0.25 |
| Urine | 0.05method1;0.25method 2 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Salbutamol | 100 |
| Salbutamol sulfate | 100 |
| Clenbuterol | 60 |
| Mapenterol | 12 |
| Tolubuterol | 6.8 |
| Terbutaline | 1.4 |
| Cimbuterol | 0.7 |
| Mabuterol | 0.5 |
| Cimaterol | 0.1 |
| Pirbuterol | <0.1 |
| Isoprenaline | <0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Rotary evaporator or nitrogen gas

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain salbutamol. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Sample Extraction

Mix 1 volume of 20X Sample Extraction Buffer with 19 volumes of distilled water.

- Preparation of 1X PBS

Mix 1 volume of the 10X PBS with 9 volumes of distilled water.

- Preparation of 1X Sample Balance Buffer

Add 10 mL of the Sample Balance Buffer I and 62 g of Sample Balance Buffer II (Powder) to a 250-mL plastic bottle. Add 190 mL of distilled water to the bottle, mix well to obtain a homogeneous solution.

### Feed

- Homogenize the feed sample with a suitable mixer.

- Add 20 mL 0.1M HCl to 2 g of the homogenized sample. Vortex for 3 minutes at maximum speed.

- Centrifuge for 10 minutes at 2,000 x g at room temperature (20 – 25C / 68 – 77F), transfer 1 mL of the supernatant to a new tube, add ~55 Lof 1 M NaOH to adjust the pH to 6-8 (the amount of 1 M NaOH could be adjusted based on different feed samples).

- Centrifuge the sample for 10 minutes at 2,000 x g. transfer 0.5 mL of the supernatant to a new tube. Add 0.5 mL 1xPBS;mix well.

- Use 50 L of the lower aqueous layer for the assay.

Note: Dilution factor: 20.

### Honey

- Add 1 g of the honey sample to 19 mL of 1X PBS, and mix well.

- mix vigorously for 10 minutes in a multi-tube vortexer or shaker

- use 50 L per well for the assay.

Note: Dilution factor: 20. If needed, the sample can be further diluted with 1X PBS.

### Meat/liver/kidney

Method I

- Remove fat from meat, liver or kidney. Homogenize the sample with a suitable mixer.

- Add 8 mL of acetonitrile and 1 mL of ethyl acetate to 3 g of the homogenized sample. Vortex for 3 minutes at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F), transfer 6 mL of the supernatant to a new tube.

- Use a rotary evaporator to dry the sample in a 60-70C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60-70C water bath.

- Add 1 mL of hexane to dissolve the sample and then add 1 mL of 1X PBS, vortex for 1 minute at maximum speed.

- Leave the tube open and heat the sample at 85C for 3 minutes ( This step can be omitted if the lower aqueous layer is enough for the ELISA assay).

- Centrifuge the sample for 5 minutes at 4,000 x g.

- Use 50 L of the lower aqueous layer for the assay.

Note: Dilution factor: 0.5.

MethodII(also Swine/Bovine Liver and Kidney )

- Remove fat from meat, liver or kidney. Homogenize the sample with a suitable mixer.

- Add 2 mL of 1X Sample Extraction Buffer to 1 g of the homogenized sample. Vortex for 3 minutes at maximum speed or rotorack for 20 minutes (e.g. VWR Tube Rotator).

- Add 3 mL of acetonitrile and 2.5 mL of 1X Sample Balance Buffer to the sample. Vortex for 1 minute at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F), transfer 0.9 mL of the supernatant to a new tube ( Avoid touching the fatty layer!).

- Use a rotary evaporatorto dry the sample in a 60C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 60C water bath.

- Add 1 mL of hexane to dissolve the ample and then add 0.6 mL of 1X PBS, vortex for 1 minute at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).discard the upper hexane layer

- Use 50 L for assay.

Note: Dilution factor: 2.

### Milk

- Take out 1 mL of milk sample, add 0.1 mL of 0.5 M HCl and 50 L of Beta-Agonist Clean Up Buffer I to all of the samples, vortex for 5 seconds, then add 50 L of Beta-Agonist Clean Up Buffer II to each sample, vortex for 30 seconds.

- Centrifuge at 2,000 x gat roomtemperature (20 –25C / 68 –77F) for 5 minutes.

- Take 100 Lof the supernatant, add 320 L of Beta-Agonist Clean Up Buffer III, mix well.

- Use 50 Lper well for the assay.

Note: Dilution factor: 5.

### Serum/Plasma

- Take out 0.5 mL of milk sample, add 0.05 mL of 0.5 M HCl and 25 L of Beta-Agonist Clean Up Buffer I to all of the samples, vortex for 5 seconds, then add 25 L of Beta-Agonist Clean Up Buffer II to each sample, vortex for 30 seconds.

- Centrifuge at 2,000 x gat room temperature (20 –25C / 68 –77F) for 5 minutes.

- Take 100 Lof the supernatant, add 320 L of Beta-Agonist Clean Up Buffer III, mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 5.

### Urine

Method I

- Centrifuge 0.5 mL of the urine sample at 4,000 x g for 5 minutes.

- Use 50 L of the supernatant per well for the assay.

Note: Dilution factor: 1. To avoid high background, we recommend diluting the sample with 1X PBS 5 fold (1:4).

Method II (Suitable for urine sample with high background)

- Take out 1 mL of sample, add 100 L of 0.5 M HCl and 50 L of Beta-Agonist Clean Up Buffer I to all of the samples, vortex for 5 seconds, then add 50 L of Beta-Agonist Clean Up Buffer II to each sample, vortex for 30 seconds.

- Centrifuge at 2,000 x gat room temperature (20 –25C / 68 –77F) for 10 minutes.

- Take 100 L of the supernatant, add 320 L of Beta-Agonist Clean Up Buffer III, mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 5.

Salbutamol ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Salbutamol Antibody #1 | 100 L | 2.4 mL |
| 1 x HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1 x Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each Salbutamol Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). 

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X HRP-Conjugated Antibody #2. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (  Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1 x Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (  Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Salbutamol Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical chloramphenicol standard curve. The sample detection and quantification limit for this kit are calculated as below.

Sample detection limit = (0.05ng/g or ppb) x (dilution factor)

Sample quantification limit = (0.15ng/g or ppb) x (dilution factor)

For example, the dilution factor for meat sample is 0.5, therefore, the detection limit for meat sample is 0.025 ng/g or ppb (0.05 ng/g x dilution factor 0.5) and the quantification limit is 0.075 ng/g or ppb (0.15 ng/g x dilution factor 0.5).

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us